## Sheet1
SITE CALIBRATION CERTIFICATE FOR | [6 empty] | Rockwell Hardness
( INDIRECT VERIFICATION )
Cali. Certi. No:- |  |  | CERT-0003 |  |  | ULR No.:- |  | CC394625000000000F |  |  |  |  | Page No. :-
Cali. Certi.Issue Date:- |  |  | 2025-04-29 |  |  | Date of Receipt : |  |  | 2025-04-29
Name of Company |  |  |  | Machine Details |  |  |  |  | Ref Std I.S:-1586-2:2018/ISO 6508-2:2015
M/s Vishal Electricals , Pune MIDC |  |  |  | Make :- |  | FIE |  |  | Calibration Procedure No. :- SUN/CPM/CP/01
|  |  |  | Model :- |  | RHT-121 |  |  | Cali. On Date :- |  |  | 2025-04-29 |  |  | Name
|  |  |  | Sr. No :- |  | FIE-RHT-008 |  |  | Date for next Cali.: |  |  | 2026-04-30 |  |  | #
|  |  |  | I.D. No :- |  | MT-003-A |  |  | (As per customer's request)
|  |  |  | Location :- |  | AT SITE |  |  | Calibrated By :- |  |  | VISHAL KHAMKAR
|  |  |  | Test Temp : |  | 28.5 °C |  |  | Condition of Item : |  |  | OK
Details of Master Test Block Used For Calibration
Sr.No | Test Block Value & Sr. No. |  |  |  | Certificate no |  |  |  |  | Calibrated by |  |  | Valid Upto
1 | HRB-1 81.3 MT-003-A |  |  |  | CERT-0003 |  |  |  |  | FRI |  |  | 2026-04-22
2 | HRA-2 81.5 MT-003-B |  |  |  | CERT-0003 |  |  |  |  | FRI |  |  | 2026-04-22
3 | HRA-3 81.6 MT-003-C |  |  |  | CERT-0003 |  |  |  |  | FRI |  |  | 2026-04-22
4 | HRD-1 81.6 MT-003-Y |  |  |  | CERT-0003 |  |  |  |  | FRI |  |  | 2027-10-29
5 | HRD-2 81.9 MT-003-R |  |  |  | CERT-0003 |  |  |  |  | FRI |  |  | 2026-10-29
6 | HRD-3 81.6 MT-003-T |  |  |  | CERT-0003 |  |  |  |  | FRI |  |  | 2027-11-29
7
8
9
10
11
12
13
14
15
16
17
18
Traceability : |  |  | Traceability to National Standards thruogh NABL Lab CC-2029
Calibration Results (All values in HR)
Sr. | Load In | Scale | Std. Test | Actual Hardness |  |  |  |  | Avg | Actual | Allow | Actual | Allow | Expanded | Allow
No. | Kgf. |  | Block Value | 1 | 2 | 3 | 4 | 5 | H | Error | Error | Repeat. | Repeat. | Uncertainty | Repeat.
1 | 15 | HR15N | 81.3 | 55 | 55 | 55 | 55 | 55 | #DIV/0! | Error | 0 | Repeat. | #DIV/0! | 1 | 1.2
2 | 15 | HR15N | 81.5 | 55 | 55 | 55 | 55 | 55 | #DIV/0! | Error | 0 | Repeat. | #DIV/0! | 1 | 1.2
3 | 15 | HR15N | 81.6 | 55 | 55 | 55 | 55 | 55 | #DIV/0! | Error | 0 | Repeat. | #DIV/0! | 1 | 1.2
4 | 45 | HR45N | 81.6 | 55 | 55 | 55 | 55 | 55 | #DIV/0! | Error | 0 | Repeat. | #DIV/0! | 1 | 1.2
5 | 45 | HR45N | 81.9 | 55 | 55 | 55 | 55 | 55 | #DIV/0! | Error | 0 | Repeat. | #DIV/0! | 1 | 1.2
6 | 45 | HR45N | 81.6 | 55 | 55 | 55 | 55 | 55 | #DIV/0! | Error | 0 | Repeat. | #DIV/0! | 1 | 1.2
7 | [8 empty] | #DIV/0! | #DIV/0! |  | 0 | #DIV/0! |  | 1.2
8 | [8 empty] | #DIV/0! | #DIV/0! |  | 0 | #DIV/0! |  | 1.2
9 | [8 empty] | #DIV/0! | #DIV/0! |  | 0 | #DIV/0! |  | 1.2
10 | [8 empty] | #DIV/0! | #DIV/0! |  | 0 | #DIV/0! |  | 0.8
11 | [8 empty] | #DIV/0! | #DIV/0! |  | 0 | #DIV/0! |  | 0.8
12 | [8 empty] | #DIV/0! | #DIV/0! |  | 0 | #DIV/0! |  | 0.8
13 | [8 empty] | #DIV/0! | #DIV/0! |  | 0 | #DIV/0!
14 | [8 empty] | #DIV/0! | #DIV/0! |  | 0 | #DIV/0!
15 | [8 empty] | #DIV/0! | #DIV/0! |  | 0 | #DIV/0!
16 | [8 empty] | #DIV/0! | #DIV/0! |  | 0 | #DIV/0! |  | 0.8
17 | [8 empty] | #DIV/0! | #DIV/0! |  | 0 | #DIV/0! |  | 0.8
18 | [8 empty] | #DIV/0! | #DIV/0! |  | 0 | #DIV/0! |  | 0.8
Other Points :- 1) Measuring Device & No. :- Dial Guage | [8 empty] | 2) Indentor & No.:-N/A
Note :- 1) Allowed Repeatability is as per Clause No. 5.3.2 of IS 1586 (Part 2):2018
2)Allowed Error is as per Clause No. 5.4.2 of IS 1586 (Part 2):2018
0
Calibrated By |  |  |  | Checked By | [5 empty] | For SUN - TECH
25/04/2023 | [11 empty] | SUN/F-34a REV : 02
• The above calibration is done at site of customer.
• The above reported expanded uncertainties are based on standard uncertainties multiplied by coverage factor K=2,
providing level of of confidence of 95.45 %.
• The calibration results reported in the certificate are valid at the time of and under the stated conditions of
measurement and is only for the calibration items as identified in the certificate.
• Calibration certificate shall not be reproduced except in full, without written approval of SUN-TECH.
"----End of Calibration Certificate----"

## uncert HR 15N
Rockwell Hardness Tester Uncertainty Calculation for Indirect Method Calibration | [8 empty] | Rockwell Hardness Tester Uncertainty Calculation for Indirect Method Calibration | [8 empty] | Rockwell Hardness Tester Uncertainty Calculation for Indirect Method Calibration | [8 empty] | Rockwell Hardness Tester Uncertainty Calculation for Indirect Method Calibration | [8 empty] | Rockwell Hardness Tester Uncertainty Calculation for Indirect Method Calibration | [9 empty] | Rockwell Hardness Tester Uncertainty Calculation for Indirect Method Calibration | [8 empty] | Rockwell Hardness Tester Uncertainty Calculation for Indirect Method Calibration | [8 empty] | Rockwell Hardness Tester Uncertainty Calculation for Indirect Method Calibration | [8 empty] | Rockwell Hardness Tester Uncertainty Calculation for Indirect Method Calibration | [8 empty] | Rockwell Hardness Tester Uncertainty Calculation for Indirect Method Calibration | [8 empty] | Rockwell Hardness Tester Uncertainty Calculation for Indirect Method Calibration | [8 empty] | Rockwell Hardness Tester Uncertainty Calculation for Indirect Method Calibration | [8 empty] | Rockwell Hardness Tester Uncertainty Calculation for Indirect Method Calibration | [8 empty] | Rockwell Hardness Tester Uncertainty Calculation for Indirect Method Calibration | [8 empty] | Rockwell Hardness Tester Uncertainty Calculation for Indirect Method Calibration | [8 empty] | Rockwell Hardness Tester Uncertainty Calculation for Indirect Method Calibration | [8 empty] | Rockwell Hardness Tester Uncertainty Calculation for Indirect Method Calibration | [8 empty] | Rockwell Hardness Tester Uncertainty Calculation for Indirect Method Calibration
| HR15N | [8 empty] | HR15N | [8 empty] | HR15N | [7 empty] | HR45N | [17 empty] | HR45N
Calibration Cert. No: |  | 0 | [6 empty] | Calibration Cert. No: |  | 0 | [6 empty] | Calibration Cert. No: |  | 0 | [6 empty] | Calibration Cert. No: |  | 0 | [6 empty] | Calibration Cert. No: |  | 0 | [7 empty] | Calibration Cert. No: |  | 0 | [6 empty] | Calibration Cert. No: |  | 0 | [6 empty] | Calibration Cert. No: |  | 0 | [6 empty] | Calibration Cert. No: |  | 0 | [6 empty] | Calibration Cert. No: |  | 0 | [6 empty] | Calibration Cert. No: |  | 0 | [6 empty] | Calibration Cert. No: |  | 0 | [6 empty] | Calibration Cert. No: |  | 0 | [6 empty] | Calibration Cert. No: |  | 0 | [6 empty] | Calibration Cert. No: |  | 0 | [6 empty] | Calibration Cert. No: |  | 0 | [6 empty] | Calibration Cert. No: |  | 0 | [6 empty] | Calibration Cert. No: |  | 0
[5 empty] | HR15N | lower | [7 empty] | HR30N | middele | [7 empty] | HR30N | HIGHER | [7 empty] | HR30N | lower | [7 empty] | HR30N | Middle | [8 empty] | HR30N | Higher | [7 empty] | HR45N | lower | [7 empty] | HR45N | Middle | [7 empty] | HR45N | Higher | [7 empty] | HRA | lower | [7 empty] | HRA | Middle | [7 empty] | HRA | Higher | [7 empty] | HRB | lower | [7 empty] | HRB | Middle | [7 empty] | HRB | Higher | [7 empty] | HRC | Lower | [7 empty] | HRC | Middle | [7 empty] | HRC | Higher
Machine Details: |  | ROCKWELL SUPERFICIAL HARDNESS TESTER | [6 empty] | Machine Details: |  | ROCKWELL SUPERFICIAL HARDNESS TESTER | [6 empty] | Machine Details: |  | ROCKWELL SUPERFICIAL HARDNESS TESTER | [6 empty] | Machine Details: |  | ROCKWELL SUPERFICIAL HARDNESS TESTER | [6 empty] | Machine Details: |  | ROCKWELL SUPERFICIAL HARDNESS TESTER | [7 empty] | Machine Details: |  | ROCKWELL SUPERFICIAL HARDNESS TESTER | [6 empty] | Machine Details: |  | ROCKWELL SUPERFICIAL HARDNESS TESTER | [6 empty] | Machine Details: |  | ROCKWELL SUPERFICIAL HARDNESS TESTER | [6 empty] | Machine Details: |  | ROCKWELL SUPERFICIAL HARDNESS TESTER | [6 empty] | Machine Details: |  | ROCKWELL SUPERFICIAL HARDNESS TESTER | [6 empty] | Machine Details: |  | ROCKWELL SUPERFICIAL HARDNESS TESTER | [6 empty] | Machine Details: |  | ROCKWELL SUPERFICIAL HARDNESS TESTER | [6 empty] | Machine Details: |  | ROCKWELL SUPERFICIAL HARDNESS TESTER | [6 empty] | Machine Details: |  | ROCKWELL SUPERFICIAL HARDNESS TESTER | [6 empty] | Machine Details: |  | ROCKWELL SUPERFICIAL HARDNESS TESTER | [6 empty] | Machine Details: |  | ROCKWELL SUPERFICIAL HARDNESS TESTER | [6 empty] | Machine Details: |  | ROCKWELL SUPERFICIAL HARDNESS TESTER | [6 empty] | Machine Details: |  | ROCKWELL SUPERFICIAL HARDNESS TESTER
|  | Make: | 0 | Model: | 0 | Sr. No.: | 0 |  |  |  | Make: | 0 | Model: | 0 | Sr. No.: | 0 |  |  |  | Make: | 0 | Model: | 0 | Sr. No.: | 0 |  |  |  | Make: | 0 | Model: | 0 | Sr. No.: | 0 |  |  |  | Make: | 0 | Model: | 0 | Sr. No.: | 0 |  |  |  |  | Make: | 0 | Model: | 0 | Sr. No.: | 0 |  |  |  | Make: | 0 | Model: | 0 | Sr. No.: | 0 |  |  |  | Make: | 0 | Model: | 0 | Sr. No.: | 0 |  |  |  | Make: | 0 | Model: | 0 | Sr. No.: | 0 |  |  |  | Make: | 0 | Model: | 0 | Sr. No.: | 0 |  |  |  | Make: | 0 | Model: | 0 | Sr. No.: | 0 |  |  |  | Make: | 0 | Model: | 0 | Sr. No.: | 0 |  |  |  | Make: | 0 | Model: | 0 | Sr. No.: | 0 |  |  |  | Make: | 0 | Model: | 0 | Sr. No.: | 0 |  |  |  | Make: | 0 | Model: | 0 | Sr. No.: | 0 |  |  |  | Make: | 0 | Model: | 0 | Sr. No.: | 0 |  |  |  | Make: | 0 | Model: | 0 | Sr. No.: | 0 |  |  |  | Make: | 0 | Model: | 0 | Sr. No.: | 0
Test block value:- |  | 0 | HR15N |  | Reprted Uncertainty:- |  | 0.76 | HR15N | Test block value:- |  | 0 | HR15N |  | Reprted Uncertainty:- |  | 0.78 | HR15N | Test block value:- |  | 0 | HR15N |  | Reprted Uncertainty:- |  | 0.76 | HR15N | Test block value:- |  | 0 | HR30N |  | Reprted Uncertainty:- |  | 0.76 | HR30N | Test block value:- |  | 0 | HR30N |  | Reprted Uncertainty:- |  | 0.76 | HR30N |  | Test block value:- |  | 0 | HR30N |  | Reprted Uncertainty:- |  | 0.76 | HR30N | Test block value:- |  | 0 | HR45N |  | Reprted Uncertainty:- |  | 0.78 | HR45N | Test block value:- |  | 0 | HR45N |  | Reprted Uncertainty:- |  | 0.76 | HR45N | Test block value:- |  | 0 | HR45N |  | Reprted Uncertainty:- |  | 0.77 | HR45N | Test block value:- |  | 0 | HRA |  | Reprted Uncertainty:- |  | 0.51 | HRA | Test block value:- |  | 0 | HRA |  | Reprted Uncertainty:- |  | 0.51 | HRA | Test block value:- |  | 0 | HRA |  | Reprted Uncertainty:- |  | 0.51 | HRA | Test block value:- |  | 0 | HRB |  | Reprted Uncertainty:- |  | 0.62 | HRB | Test block value:- |  | 0 | HRB |  | Reprted Uncertainty:- |  | 0.6 | HRB | Test block value:- |  | 0 | HRB |  | Reprted Uncertainty:- |  | 0.55 | HRB | Test block value:- |  | 0 | HRC |  | Reprted Uncertainty:- |  | 0.52 | HRC | Test block value:- |  | 0 | HRC |  | Reprted Uncertainty:- |  | 0.51 | HRC | Test block value:- |  | 0 | HRC |  | Reprted Uncertainty:- |  | 0.51 | HRC
Type:- A | [8 empty] | Type:- A | [8 empty] | Type:- A | [8 empty] | Type:- A | [8 empty] | Type:- A | [9 empty] | Type:- A | [8 empty] | Type:- A | [8 empty] | Type:- A | [8 empty] | Type:- A | [8 empty] | Type:- A | [8 empty] | Type:- A | [8 empty] | Type:- A | [8 empty] | Type:- A | [8 empty] | Type:- A | [8 empty] | Type:- A | [8 empty] | Type:- A | [8 empty] | Type:- A | [8 empty] | Type:- A
X₁ | X | (X₁ - X) | (X₁ - X)² |  | Standard Deviation |  |  |  | X₁ | X | (X₁ - X) | (X₁ - X)² |  | Standard Deviation |  |  |  | X₁ | X | (X₁ - X) | (X₁ - X)² |  | Standard Deviation |  |  |  | X₁ | X | (X₁ - X) | (X₁ - X)² |  | Standard Deviation |  |  |  | X₁ | X | (X₁ - X) | (X₁ - X)² |  | Standard Deviation |  |  |  |  | X₁ | X | (X₁ - X) | (X₁ - X)² |  | Standard Deviation |  |  |  | X₁ | X | (X₁ - X) | (X₁ - X)² |  | Standard Deviation |  |  |  | X₁ | X | (X₁ - X) | (X₁ - X)² |  | Standard Deviation |  |  |  | X₁ | X | (X₁ - X) | (X₁ - X)² |  | Standard Deviation |  |  |  | X₁ | X | (X₁ - X) | (X₁ - X)² |  | Standard Deviation |  |  |  | X₁ | X | (X₁ - X) | (X₁ - X)² |  | Standard Deviation |  |  |  | X₁ | X | (X₁ - X) | (X₁ - X)² |  | Standard Deviation |  |  |  | X₁ | X | (X₁ - X) | (X₁ - X)² |  | Standard Deviation |  |  |  | X₁ | X | (X₁ - X) | (X₁ - X)² |  | Standard Deviation |  |  |  | X₁ | X | (X₁ - X) | (X₁ - X)² |  | Standard Deviation |  |  |  | X₁ | X | (X₁ - X) | (X₁ - X)² |  | Standard Deviation |  |  |  | X₁ | X | (X₁ - X) | (X₁ - X)² |  | Standard Deviation |  |  |  | X₁ | X | (X₁ - X) | (X₁ - X)² |  | Standard Deviation
0 | 0 | 0 | 0 | [5 empty] | 0 | 0 | 0 | 0 | [5 empty] | 0 | 0 | 0 | 0 | [5 empty] | 0 | 0 | 0 | 0 | [5 empty] | 0 | 0 | 0 | 0 | [6 empty] | 0 | 0 | 0 | 0 | [5 empty] | 0 | 0 | 0 | 0 | [5 empty] | 0 | 0 | 0 | 0 | [5 empty] | 0 | 0 | 0 | 0 | [5 empty] | 0 | 0 | 0 | 0 | [5 empty] | 0 | 0 | 0 | 0 | [5 empty] | 0 | 0 | 0 | 0 | [5 empty] | 0 | 0 | 0 | 0 | [5 empty] | 0 | 0 | 0 | 0 | [5 empty] | 0 | 0 | 0 | 0 | [5 empty] | 0 | 0 | 0 | 0 | [5 empty] | 0 | 0 | 0 | 0 | [5 empty] | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | б = | ∑ (X₁ - X)² |  |  |  | 0 | 0 | 0 | 0 | б = | ∑ (X₁ - X)² |  |  |  | 0 | 0 | 0 | 0 | б = | ∑ (X₁ - X)² |  |  |  | 0 | 0 | 0 | 0 | б = | ∑ (X₁ - X)² |  |  |  | 0 | 0 | 0 | 0 | б = | ∑ (X₁ - X)² |  |  |  |  | 0 | 0 | 0 | 0 | б = | ∑ (X₁ - X)² |  |  |  | 0 | 0 | 0 | 0 | б = | ∑ (X₁ - X)² |  |  |  | 0 | 0 | 0 | 0 | б = | ∑ (X₁ - X)² |  |  |  | 0 | 0 | 0 | 0 | б = | ∑ (X₁ - X)² |  |  |  | 0 | 0 | 0 | 0 | б = | ∑ (X₁ - X)² |  |  |  | 0 | 0 | 0 | 0 | б = | ∑ (X₁ - X)² |  |  |  | 0 | 0 | 0 | 0 | б = | ∑ (X₁ - X)² |  |  |  | 0 | 0 | 0 | 0 | б = | ∑ (X₁ - X)² |  |  |  | 0 | 0 | 0 | 0 | б = | ∑ (X₁ - X)² |  |  |  | 0 | 0 | 0 | 0 | б = | ∑ (X₁ - X)² |  |  |  | 0 | 0 | 0 | 0 | б = | ∑ (X₁ - X)² |  |  |  | 0 | 0 | 0 | 0 | б = | ∑ (X₁ - X)² |  |  |  | 0 | 0 | 0 | 0 | б = | ∑ (X₁ - X)²
0 | 0 | 0 | 0 |  | n - 1 | 4 |  |  | 0 | 0 | 0 | 0 |  | n - 1 | 4 |  |  | 0 | 0 | 0 | 0 |  | n - 1 | 4 |  |  | 0 | 0 | 0 | 0 |  | n - 1 | 4 |  |  | 0 | 0 | 0 | 0 |  | n - 1 | 4 |  |  |  | 0 | 0 | 0 | 0 |  | n - 1 | 4 |  |  | 0 | 0 | 0 | 0 |  | n - 1 | 4 |  |  | 0 | 0 | 0 | 0 |  | n - 1 | 4 |  |  | 0 | 0 | 0 | 0 |  | n - 1 | 4 |  |  | 0 | 0 | 0 | 0 |  | n - 1 | 4 |  |  | 0 | 0 | 0 | 0 |  | n - 1 | 4 |  |  | 0 | 0 | 0 | 0 |  | n - 1 | 4 |  |  | 0 | 0 | 0 | 0 |  | n - 1 | 4 |  |  | 0 | 0 | 0 | 0 |  | n - 1 | 4 |  |  | 0 | 0 | 0 | 0 |  | n - 1 | 4 |  |  | 0 | 0 | 0 | 0 |  | n - 1 | 4 |  |  | 0 | 0 | 0 | 0 |  | n - 1 | 4 |  |  | 0 | 0 | 0 | 0 |  | n - 1 | 4
0 | 0 | 0 | 0 | [5 empty] | 0 | 0 | 0 | 0 | [5 empty] | 0 | 0 | 0 | 0 | [5 empty] | 0 | 0 | 0 | 0 | [5 empty] | 0 | 0 | 0 | 0 | [6 empty] | 0 | 0 | 0 | 0 | [5 empty] | 0 | 0 | 0 | 0 | [5 empty] | 0 | 0 | 0 | 0 | [5 empty] | 0 | 0 | 0 | 0 | [5 empty] | 0 | 0 | 0 | 0 | [5 empty] | 0 | 0 | 0 | 0 | [5 empty] | 0 | 0 | 0 | 0 | [5 empty] | 0 | 0 | 0 | 0 | [5 empty] | 0 | 0 | 0 | 0 | [5 empty] | 0 | 0 | 0 | 0 | [5 empty] | 0 | 0 | 0 | 0 | [5 empty] | 0 | 0 | 0 | 0 | [5 empty] | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | б = | 0 |  |  |  | 0 | 0 | 0 | 0 | б = | 0 |  |  |  | 0 | 0 | 0 | 0 | б = | 0 |  |  |  | 0 | 0 | 0 | 0 | б = | 0 |  |  |  | 0 | 0 | 0 | 0 | б = | 0 |  |  |  |  | 0 | 0 | 0 | 0 | б = | 0 |  |  |  | 0 | 0 | 0 | 0 | б = | 0 |  |  |  | 0 | 0 | 0 | 0 | б = | 0 |  |  |  | 0 | 0 | 0 | 0 | б = | 0 |  |  |  | 0 | 0 | 0 | 0 | б = | 0 |  |  |  | 0 | 0 | 0 | 0 | б = | 0 |  |  |  | 0 | 0 | 0 | 0 | б = | 0 |  |  |  | 0 | 0 | 0 | 0 | б = | 0 |  |  |  | 0 | 0 | 0 | 0 | б = | 0 |  |  |  | 0 | 0 | 0 | 0 | б = | 0 |  |  |  | 0 | 0 | 0 | 0 | б = | 0 |  |  |  | 0 | 0 | 0 | 0 | б = | 0 |  |  |  | 0 | 0 | 0 | 0 | б = | 0
0 |  | Sum = | 0 | [5 empty] | 0 |  | Sum = | 0 | [5 empty] | 0 |  | Sum = | 0 | [5 empty] | 0 |  | Sum = | 0 | [5 empty] | 0 |  | Sum = | 0 | [6 empty] | 0 |  | Sum = | 0 | [5 empty] | 0 |  | Sum = | 0 | [5 empty] | 0 |  | Sum = | 0 | [5 empty] | 0 |  | Sum = | 0 | [5 empty] | 0 |  | Sum = | 0 | [5 empty] | 0 |  | Sum = | 0 | [5 empty] | 0 |  | Sum = | 0 | [5 empty] | 0 |  | Sum = | 0 | [5 empty] | 0 |  | Sum = | 0 | [5 empty] | 0 |  | Sum = | 0 | [5 empty] | 0 |  | Sum = | 0 | [5 empty] | 0 |  | Sum = | 0 | [5 empty] | 0 |  | Sum = | 0
|  |  |  | б = | 0 | [7 empty] | б = | 0 | [7 empty] | б = | 0 | [7 empty] | б = | 0 | [7 empty] | б = | 0 | [8 empty] | б = | 0 | [7 empty] | б = | 0 | [7 empty] | б = | 0 | [7 empty] | б = | 0 | [7 empty] | б = | 0 | [7 empty] | б = | 0 | [7 empty] | б = | 0 | [7 empty] | б = | 0 | [7 empty] | б = | 0 | [7 empty] | б = | 0 | [7 empty] | б = | 0 | [7 empty] | б = | 0 | [7 empty] | б = | 0
Standard Deviation: |  | (б/√n)*t = | 0 |  | where t = 1.14 for 5 readings |  |  |  | Standard Deviation: |  | (б/√n)*t = | 0 |  | where t = 1.14 for 5 readings |  |  |  | Standard Deviation: |  | (б/√n)*t = | 0 |  | where t = 1.14 for 5 readings |  |  |  | Standard Deviation: |  | (б/√n)*t = | 0 |  | where t = 1.14 for 5 readings |  |  |  | Standard Deviation: |  | (б/√n)*t = | 0 |  | where t = 1.14 for 5 readings |  |  |  |  | Standard Deviation: |  | (б/√n)*t = | 0 |  | where t = 1.14 for 5 readings |  |  |  | Standard Deviation: |  | (б/√n)*t = | 0 |  | where t = 1.14 for 5 readings |  |  |  | Standard Deviation: |  | (б/√n)*t = | 0 |  | where t = 1.14 for 5 readings |  |  |  | Standard Deviation: |  | (б/√n)*t = | 0 |  | where t = 1.14 for 5 readings |  |  |  | Standard Deviation: |  | (б/√n)*t = | 0 |  | where t = 1.14 for 5 readings |  |  |  | Standard Deviation: |  | (б/√n)*t = | 0 |  | where t = 1.14 for 5 readings |  |  |  | Standard Deviation: |  | (б/√n)*t = | 0 |  | where t = 1.14 for 5 readings |  |  |  | Standard Deviation: |  | (б/√n)*t = | 0 |  | where t = 1.14 for 5 readings |  |  |  | Standard Deviation: |  | (б/√n)*t = | 0 |  | where t = 1.14 for 5 readings |  |  |  | Standard Deviation: |  | (б/√n)*t = | 0 |  | where t = 1.14 for 5 readings |  |  |  | Standard Deviation: |  | (б/√n)*t = | 0 |  | where t = 1.14 for 5 readings |  |  |  | Standard Deviation: |  | (б/√n)*t = | 0 |  | where t = 1.14 for 5 readings |  |  |  | Standard Deviation: |  | (б/√n)*t = | 0 |  | where t = 1.14 for 5 readings
Type:- B | [8 empty] | Type:- B | [8 empty] | Type:- B | [8 empty] | Type:- B | [8 empty] | Type:- B | [9 empty] | Type:- B | [8 empty] | Type:- B | [8 empty] | Type:- B | [8 empty] | Type:- B | [8 empty] | Type:- B | [8 empty] | Type:- B | [8 empty] | Type:- B | [8 empty] | Type:- B | [8 empty] | Type:- B | [8 empty] | Type:- B | [8 empty] | Type:- B | [8 empty] | Type:- B | [8 empty] | Type:- B
Source of Uncertainty | [8 empty] | Source of Uncertainty | [8 empty] | Source of Uncertainty | [8 empty] | Source of Uncertainty | [8 empty] | Source of Uncertainty | [9 empty] | Source of Uncertainty | [8 empty] | Source of Uncertainty | [8 empty] | Source of Uncertainty | [8 empty] | Source of Uncertainty | [8 empty] | Source of Uncertainty | [8 empty] | Source of Uncertainty | [8 empty] | Source of Uncertainty | [8 empty] | Source of Uncertainty | [8 empty] | Source of Uncertainty | [8 empty] | Source of Uncertainty | [8 empty] | Source of Uncertainty | [8 empty] | Source of Uncertainty | [8 empty] | Source of Uncertainty
S.No. | Source-Reference-Value |  |  | Distribution | DOF | Calculation |  | Value | S.No. | Source-Reference-Value |  |  | Distribution | DOF | Calculation |  | Value | S.No. | Source-Reference-Value |  |  | Distribution | DOF | Calculation |  | Value | S.No. | Source-Reference-Value |  |  | Distribution | DOF | Calculation |  | Value | S.No. | Source-Reference-Value |  |  | Distribution | DOF | Calculation |  | Value |  | S.No. | Source-Reference-Value |  |  | Distribution | DOF | Calculation |  | Value | S.No. | Source-Reference-Value |  |  | Distribution | DOF | Calculation |  | Value | S.No. | Source-Reference-Value |  |  | Distribution | DOF | Calculation |  | Value | S.No. | Source-Reference-Value |  |  | Distribution | DOF | Calculation |  | Value | S.No. | Source-Reference-Value |  |  | Distribution | DOF | Calculation |  | Value | S.No. | Source-Reference-Value |  |  | Distribution | DOF | Calculation |  | Value | S.No. | Source-Reference-Value |  |  | Distribution | DOF | Calculation |  | Value | S.No. | Source-Reference-Value |  |  | Distribution | DOF | Calculation |  | Value | S.No. | Source-Reference-Value |  |  | Distribution | DOF | Calculation |  | Value | S.No. | Source-Reference-Value |  |  | Distribution | DOF | Calculation |  | Value | S.No. | Source-Reference-Value |  |  | Distribution | DOF | Calculation |  | Value | S.No. | Source-Reference-Value |  |  | Distribution | DOF | Calculation |  | Value | S.No. | Source-Reference-Value |  |  | Distribution | DOF | Calculation |  | Value
|  |  |  | Type | [8 empty] | Type | [8 empty] | Type | [8 empty] | Type | [8 empty] | Type | [9 empty] | Type | [8 empty] | Type | [8 empty] | Type | [8 empty] | Type | [8 empty] | Type | [8 empty] | Type | [8 empty] | Type | [8 empty] | Type | [8 empty] | Type | [8 empty] | Type | [8 empty] | Type | [8 empty] | Type | [8 empty] | Type
UA | Standard Deviation A Type |  |  | Normal |  |  |  |  | UA | Standard Deviation A Type |  |  | Normal |  |  |  |  | UA | Standard Deviation A Type |  |  | Normal |  |  |  |  | UA | Standard Deviation A Type |  |  | Normal |  |  |  |  | UA | Standard Deviation A Type |  |  | Normal | [5 empty] | UA | Standard Deviation A Type |  |  | Normal |  |  |  |  | UA | Standard Deviation A Type |  |  | Normal |  |  |  |  | UA | Standard Deviation A Type |  |  | Normal |  |  |  |  | UA | Standard Deviation A Type |  |  | Normal |  |  |  |  | UA | Standard Deviation A Type |  |  | Normal |  |  |  |  | UA | Standard Deviation A Type |  |  | Normal |  |  |  |  | UA | Standard Deviation A Type |  |  | Normal |  |  |  |  | UA | Standard Deviation A Type |  |  | Normal |  |  |  |  | UA | Standard Deviation A Type |  |  | Normal |  |  |  |  | UA | Standard Deviation A Type |  |  | Normal |  |  |  |  | UA | Standard Deviation A Type |  |  | Normal |  |  |  |  | UA | Standard Deviation A Type |  |  | Normal |  |  |  |  | UA | Standard Deviation A Type |  |  | Normal
|  |  |  | Tupe A | 4 | 0 | /2.24 | 0 |  |  |  |  | Tupe A | 4 | 0 | /2.24 | 0 |  |  |  |  | Tupe A | 4 | 0 | /2.24 | 0 |  |  |  |  | Tupe A | 4 | 0 | /2.24 | 0 |  |  |  |  | Tupe A | 4 | 0 | /2.24 | 0 | [5 empty] | Tupe A | 4 | 0 | /2.24 | 0 |  |  |  |  | Tupe A | 4 | 0 | /2.24 | 0 |  |  |  |  | Tupe A | 4 | 0 | /2.24 | 0 |  |  |  |  | Tupe A | 4 | 0 | /2.24 | 0 |  |  |  |  | Tupe A | 4 | 0 | /2.24 | 0 |  |  |  |  | Tupe A | 4 | 0 | /2.24 | 0 |  |  |  |  | Tupe A | 4 | 0 | /2.24 | 0 |  |  |  |  | Tupe A | 4 | 0 | /2.24 | 0 |  |  |  |  | Tupe A | 4 | 0 | /2.24 | 0 |  |  |  |  | Tupe A | 4 | 0 | /2.24 | 0 |  |  |  |  | Tupe A | 4 | 0 | /2.24 | 0 |  |  |  |  | Tupe A | 4 | 0 | /2.24 | 0 |  |  |  |  | Tupe A | 4 | 0 | /2.24 | 0
UB1 | Reported Uncertainty for Test |  |  | Normal |  |  |  |  | UB1 | Reported Uncertainty for Test |  |  | Normal |  |  |  |  | UB1 | Reported Uncertainty for Test |  |  | Normal |  |  |  |  | UB1 | Reported Uncertainty for Test |  |  | Normal |  |  |  |  | UB1 | Reported Uncertainty for Test |  |  | Normal | [5 empty] | UB1 | Reported Uncertainty for Test |  |  | Normal |  |  |  |  | UB1 | Reported Uncertainty for Test |  |  | Normal |  |  |  |  | UB1 | Reported Uncertainty for Test |  |  | Normal |  |  |  |  | UB1 | Reported Uncertainty for Test |  |  | Normal |  |  |  |  | UB1 | Reported Uncertainty for Test |  |  | Normal |  |  |  |  | UB1 | Reported Uncertainty for Test |  |  | Normal |  |  |  |  | UB1 | Reported Uncertainty for Test |  |  | Normal |  |  |  |  | UB1 | Reported Uncertainty for Test |  |  | Normal |  |  |  |  | UB1 | Reported Uncertainty for Test |  |  | Normal |  |  |  |  | UB1 | Reported Uncertainty for Test |  |  | Normal |  |  |  |  | UB1 | Reported Uncertainty for Test |  |  | Normal |  |  |  |  | UB1 | Reported Uncertainty for Test |  |  | Normal |  |  |  |  | UB1 | Reported Uncertainty for Test |  |  | Normal
| Block | 0.76 |  | Type B |  | 0.59/2 |  | 0.38 |  | Block | 0.78 |  | Type B |  | 0.48/2 |  | 0.39 |  | Block | 0.76 |  | Type B |  | 0.47/2 |  | 0.38 |  | Block | 0.76 |  | Type B |  | 0.59/2 |  | 0.38 |  | Block | 0.76 |  | Type B |  | 0.59/2 |  | 0.38 |  |  | Block | 0.76 |  | Type B |  | 0.59/2 |  | 0.38 |  | Block | 0.78 |  | Type B |  | 0.59/2 |  | 0.39 |  | Block | 0.76 |  | Type B |  | 0.59/2 |  | 0.38 |  | Block | 0.77 |  | Type B |  | 0.59/2 |  | 0.385 |  | Block | 0.51 |  | Type B |  | 0.59/2 |  | 0.255 |  | Block | 0.51 |  | Type B |  | 0.59/2 |  | 0.255 |  | Block | 0.51 |  | Type B |  | 0.59/2 |  | 0.255 |  | Block | 0.62 |  | Type B |  | 0.59/2 |  | 0.31 |  | Block | 0.6 |  | Type B |  | 0.59/2 |  | 0.3 |  | Block | 0.55 |  | Type B |  | 0.59/2 |  | 0.275 |  | Block | 0.52 |  | Type B |  | 0.59/2 |  | 0.26 |  | Block | 0.51 |  | Type B |  | 0.59/2 |  | 0.255 |  | Block | 0.51 |  | Type B |  | 0.59/2 |  | 0.255
UB2 | Uncertainty due to resolution of |  |  | Rectangular |  |  |  |  | UB2 | Uncertainty due to resolution of |  |  | Rectangular |  |  |  |  | UB2 | Uncertainty due to resolution of |  |  | Rectangular |  |  |  |  | UB2 | Uncertainty due to resolution of |  |  | Rectangular |  |  |  |  | UB2 | Uncertainty due to resolution of |  |  | Rectangular | [5 empty] | UB2 | Uncertainty due to resolution of |  |  | Rectangular |  |  |  |  | UB2 | Uncertainty due to resolution of |  |  | Rectangular |  |  |  |  | UB2 | Uncertainty due to resolution of |  |  | Rectangular |  |  |  |  | UB2 | Uncertainty due to resolution of |  |  | Rectangular |  |  |  |  | UB2 | Uncertainty due to resolution of |  |  | Rectangular |  |  |  |  | UB2 | Uncertainty due to resolution of |  |  | Rectangular |  |  |  |  | UB2 | Uncertainty due to resolution of |  |  | Rectangular |  |  |  |  | UB2 | Uncertainty due to resolution of |  |  | Rectangular |  |  |  |  | UB2 | Uncertainty due to resolution of |  |  | Rectangular |  |  |  |  | UB2 | Uncertainty due to resolution of |  |  | Rectangular |  |  |  |  | UB2 | Uncertainty due to resolution of |  |  | Rectangular |  |  |  |  | UB2 | Uncertainty due to resolution of |  |  | Rectangular |  |  |  |  | UB2 | Uncertainty due to resolution of |  |  | Rectangular
| DUC 0.2 |  |  | Type B |  | 0.2/2/1.73 |  | 0.057803 |  | DUC 0.2 |  |  | Type B |  | 0.2/2/1.73 |  | 0.057803 |  | DUC 0.2 |  |  | Type B |  | 0.2/2/1.73 |  | 0.057803 |  | DUC 0.2 |  |  | Type B |  | 0.2/2/1.73 |  | 0.057803 |  | DUC 0.2 |  |  | Type B |  | 0.2/2/1.73 |  | 0.057803 |  |  | DUC 0.2 |  |  | Type B |  | 0.2/2/1.73 |  | 0.057803 |  | DUC 0.2 |  |  | Type B |  | 0.2/2/1.73 |  | 0.057803 |  | DUC 0.2 |  |  | Type B |  | 0.2/2/1.73 |  | 0.057803 |  | DUC 0.2 |  |  | Type B |  | 0.2/2/1.73 |  | 0.057803 |  | DUC 0.2 |  |  | Type B |  | 0.2/2/1.73 |  | 0.057803 |  | DUC 0.2 |  |  | Type B |  | 0.2/2/1.73 |  | 0.057803 |  | DUC 0.2 |  |  | Type B |  | 0.2/2/1.73 |  | 0.057803 |  | DUC 0.2 |  |  | Type B |  | 0.2/2/1.73 |  | 0.057803 |  | DUC 0.2 |  |  | Type B |  | 0.2/2/1.73 |  | 0.057803 |  | DUC 0.2 |  |  | Type B |  | 0.2/2/1.73 |  | 0.057803 |  | DUC 0.2 |  |  | Type B |  | 0.2/2/1.73 |  | 0.057803 |  | DUC 0.2 |  |  | Type B |  | 0.2/2/1.73 |  | 0.057803 |  | DUC 0.2 |  |  | Type B |  | 0.2/2/1.73 |  | 0.057803
| UA² + UB1² + UB2² | [8 empty] | UA² + UB1² + UB2² | [8 empty] | UA² + UB1² + UB2² | [8 empty] | UA² + UB1² + UB2² | [8 empty] | UA² + UB1² + UB2² | [9 empty] | UA² + UB1² + UB2² | [8 empty] | UA² + UB1² + UB2² | [8 empty] | UA² + UB1² + UB2² | [8 empty] | UA² + UB1² + UB2² | [8 empty] | UA² + UB1² + UB2² | [8 empty] | UA² + UB1² + UB2² | [8 empty] | UA² + UB1² + UB2² | [8 empty] | UA² + UB1² + UB2² | [8 empty] | UA² + UB1² + UB2² | [8 empty] | UA² + UB1² + UB2² | [8 empty] | UA² + UB1² + UB2² | [8 empty] | UA² + UB1² + UB2² | [8 empty] | UA² + UB1² + UB2²
| 0.14774 | [8 empty] | 0.15544 | [8 empty] | 0.14774 | [8 empty] | 0.14774 | [8 empty] | 0.14774 | [9 empty] | 0.14774 | [8 empty] | 0.15544 | [8 empty] | 0.14774 | [8 empty] | 0.15157 | [8 empty] | 0.068366 | [8 empty] | 0.068366 | [8 empty] | 0.068366 | [8 empty] | 0.099441 | [8 empty] | 0.093341 | [8 empty] | 0.078966 | [8 empty] | 0.070941 | [8 empty] | 0.068366 | [8 empty] | 0.068366
Combined Uncertainty Uc :- |  |  | 0.38437 | [5 empty] | Combined Uncertainty Uc :- |  |  | 0.39426 | [5 empty] | Combined Uncertainty Uc :- |  |  | 0.38437 | [5 empty] | Combined Uncertainty Uc :- |  |  | 0.38437 | [5 empty] | Combined Uncertainty Uc :- |  |  | 0.38437 | [6 empty] | Combined Uncertainty Uc :- |  |  | 0.38437 | [5 empty] | Combined Uncertainty Uc :- |  |  | 0.39426 | [5 empty] | Combined Uncertainty Uc :- |  |  | 0.38437 | [5 empty] | Combined Uncertainty Uc :- |  |  | 0.38932 | [5 empty] | Combined Uncertainty Uc :- |  |  | 0.26147 | [5 empty] | Combined Uncertainty Uc :- |  |  | 0.26147 | [5 empty] | Combined Uncertainty Uc :- |  |  | 0.26147 | [5 empty] | Combined Uncertainty Uc :- |  |  | 0.31534 | [5 empty] | Combined Uncertainty Uc :- |  |  | 0.30552 | [5 empty] | Combined Uncertainty Uc :- |  |  | 0.28101 | [5 empty] | Combined Uncertainty Uc :- |  |  | 0.26635 | [5 empty] | Combined Uncertainty Uc :- |  |  | 0.26147 | [5 empty] | Combined Uncertainty Uc :- |  |  | 0.26147
Degree of Freedom :- |  | (UC4 / UA4)* (n-1) = |  | #DIV/0! |  |  |  |  | Degree of Freedom :- |  | (UC4 / UA4)* (n-1) = |  | #DIV/0! |  |  |  |  | Degree of Freedom :- |  | (UC4 / UA4)* (n-1) = |  | #DIV/0! |  |  |  |  | Degree of Freedom :- |  | (UC4 / UA4)* (n-1) = |  | #DIV/0! |  |  |  |  | Degree of Freedom :- |  | (UC4 / UA4)* (n-1) = |  | #DIV/0! | [5 empty] | Degree of Freedom :- |  | (UC4 / UA4)* (n-1) = |  | #DIV/0! |  |  |  |  | Degree of Freedom :- |  | (UC4 / UA4)* (n-1) = |  | #DIV/0! |  |  |  |  | Degree of Freedom :- |  | (UC4 / UA4)* (n-1) = |  | #DIV/0! |  |  |  |  | Degree of Freedom :- |  | (UC4 / UA4)* (n-1) = |  | #DIV/0! |  |  |  |  | Degree of Freedom :- |  | (UC4 / UA4)* (n-1) = |  | #DIV/0! |  |  |  |  | Degree of Freedom :- |  | (UC4 / UA4)* (n-1) = |  | #DIV/0! |  |  |  |  | Degree of Freedom :- |  | (UC4 / UA4)* (n-1) = |  | #DIV/0! |  |  |  |  | Degree of Freedom :- |  | (UC4 / UA4)* (n-1) = |  | #DIV/0! |  |  |  |  | Degree of Freedom :- |  | (UC4 / UA4)* (n-1) = |  | #DIV/0! |  |  |  |  | Degree of Freedom :- |  | (UC4 / UA4)* (n-1) = |  | #DIV/0! |  |  |  |  | Degree of Freedom :- |  | (UC4 / UA4)* (n-1) = |  | #DIV/0! |  |  |  |  | Degree of Freedom :- |  | (UC4 / UA4)* (n-1) = |  | #DIV/0! |  |  |  |  | Degree of Freedom :- |  | (UC4 / UA4)* (n-1) = |  | #DIV/0!
Coverage Factor : K = |  | #DIV/0! | [6 empty] | Coverage Factor : K = |  | #DIV/0! | [6 empty] | Coverage Factor : K = |  | #DIV/0! | [6 empty] | Coverage Factor : K = |  | #DIV/0! | [6 empty] | Coverage Factor : K = |  | #DIV/0! | [7 empty] | Coverage Factor : K = |  | #DIV/0! | [6 empty] | Coverage Factor : K = |  | #DIV/0! | [6 empty] | Coverage Factor : K = |  | #DIV/0! | [6 empty] | Coverage Factor : K = |  | #DIV/0! | [6 empty] | Coverage Factor : K = |  | #DIV/0! | [6 empty] | Coverage Factor : K = |  | #DIV/0! | [6 empty] | Coverage Factor : K = |  | #DIV/0! | [6 empty] | Coverage Factor : K = |  | #DIV/0! | [6 empty] | Coverage Factor : K = |  | #DIV/0! | [6 empty] | Coverage Factor : K = |  | #DIV/0! | [6 empty] | Coverage Factor : K = |  | #DIV/0! | [6 empty] | Coverage Factor : K = |  | #DIV/0! | [6 empty] | Coverage Factor : K = |  | #DIV/0!
Expanded Uncertainty UcxK :- |  |  | #DIV/0! | HR 15N | < | 1 HR 15N | As per scope |  | Expanded Uncertainty UcxK :- |  |  | #DIV/0! | HR 15N | < | 1 HR 15N | As per scope |  | Expanded Uncertainty UcxK :- |  |  | #DIV/0! | HR 15N | < | 1 HR 15N | As per scope |  | Expanded Uncertainty UcxK :- |  |  | #DIV/0! | HR 30N | < | 1 HR 30N | As per scope |  | Expanded Uncertainty UcxK :- |  |  | #DIV/0! | HR 30N | < | 1 HR 30N | As per scope |  |  | Expanded Uncertainty UcxK :- |  |  | #DIV/0! | HR 30N | < | 1 HR 30N | As per scope |  | Expanded Uncertainty UcxK :- |  |  | #DIV/0! | HR45N | < | 1 HR 45N | As per scope |  | Expanded Uncertainty UcxK :- |  |  | #DIV/0! | HR45N | < | 1 HR 45N | As per scope |  | Expanded Uncertainty UcxK :- |  |  | #DIV/0! | HR45N | < | 1 HR 45N | As per scope |  | Expanded Uncertainty UcxK :- |  |  | #DIV/0! | HRA | < | 1 HRA | As per scope |  | Expanded Uncertainty UcxK :- |  |  | #DIV/0! | HRA | < | 1 HRA | As per scope |  | Expanded Uncertainty UcxK :- |  |  | #DIV/0! | HRA | < | 1 HRA | As per scope |  | Expanded Uncertainty UcxK :- |  |  | #DIV/0! | HRB | < | 1 HRB | As per scope |  | Expanded Uncertainty UcxK :- |  |  | #DIV/0! | HRB | < | 1 HRB | As per scope |  | Expanded Uncertainty UcxK :- |  |  | #DIV/0! | HRB | < | 1 HRB | As per scope |  | Expanded Uncertainty UcxK :- |  |  | #DIV/0! | HRC | < | 1 HRC | As per scope |  | Expanded Uncertainty UcxK :- |  |  | #DIV/0! | HRC | < | 1 HRC | As per scope |  | Expanded Uncertainty UcxK :- |  |  | #DIV/0! | HRC | < | 1 HRC | As per scope
| Reported Uncertainty= |  | 1 | HR15N | [5 empty] | Reported Uncertainty= |  | 1 | HR15N | [5 empty] | Reported Uncertainty= |  | 1 | HR15N | [5 empty] | HR45N |  | 1 | HR 30N | [5 empty] | Reported Uncertainty= |  | 1 | HR 30N | [6 empty] | HR45N |  | 1 | HR 30N | [5 empty] | Reported Uncertainty= |  | 1 | HR45N | [5 empty] | Reported Uncertainty= |  | 1 | HR45N | [5 empty] | Reported Uncertainty= |  | 1 | HR45N | [5 empty] | Reported Uncertainty= |  | 1 | HRA | [5 empty] | Reported Uncertainty= |  | 1 | HRA | [5 empty] | Reported Uncertainty= |  | 1 | HRA | [5 empty] | Reported Uncertainty= |  | 1 | HRB | [5 empty] | Reported Uncertainty= |  | 1 | HRB | [5 empty] | Reported Uncertainty= |  | 1 | HRB | [5 empty] | Reported Uncertainty= |  | 1 | HRC | [5 empty] | Reported Uncertainty= |  | 1 | HRC | [5 empty] | Reported Uncertainty= |  | 1 | HRC

## Sheet3
Rockwell Hardness Tester Uncertainty Calculation for Indirect Method Calibration
Calibration Cert. No: |  | SUN/12-13/121/196
Machine Details: |  | Rockwell Hardness Tester
|  | Make: | SAROJ | Model: | RAS | Sr. No.: | 26/24
Test block value:- |  | 60.42 |  |  | Reprted Uncertainty:- |  | 0.33
Type:- A
X₁ | X | (X₁ - X) | (X₁ - X)² |  | Standard Deviation
60.2 | 60.52 | -0.32 | 0.1024
60.8 | 60.52 | 0.28 | 0.0784 | б = | ∑ (X₁ - X)²
60.8 | 60.52 | 0.28 | 0.0784 |  | n - 1 | 4
60.4 | 60.52 | -0.12 | 0.0144
60.4 | 60.52 | -0.12 | 0.0144 | б = | 0.072
60.52 |  | Sum = | 0.288
|  |  |  | б = | 0.26833
Standard Deviation: |  | б/√n = | 0.11979
Type:- B
Source of Uncertainty
S.No. | Source-Reference-Value |  |  | Distribution | DOF | Calculation |  | Value
Type
UA | Standard Deviation A Type |  |  | Normal
|  |  |  | Tupe A | 4 | 0.268/2.24 |  | 0.11979
UB1 | Reported Uncertainty for Test |  |  | Normal
| Block | 0.33 |  | Type B | ∞ | 0.33/2 |  | 0.165
UB2 | Uncertainty due to resolution of |  |  | Rectangular
| DUC 0.2 |  |  | Type B | ∞ | 0.2/2/1.73 |  | 0.057803
UA² + UB1² + UB2²
0.044916
Combined Uncertainty Uc :- |  |  | 0.21193
Degree of Freedom :- |  |  | 39.1908
Coverage Factor : K = 2
Expanded Uncertainty UcxK :- |  |  | 0.42387 | HRA